AeroMedic: Project Master Blueprint & Systems Engineering Specification

Comprehensive Planning and Execution Document

Utilizing Systems Engineering V-Model Methodology

System Classification: Cyber-Physical Emergency System for Automated Medical Payload Delivery.

Developed By: Norah Alhussain

Academic Major: Internet of Things (IoT)

Institution: Princess Nourah University

1. Systems Engineering Management Plan (SEMP)

This plan serves as the governing constitution for the AeroMedic project. It establishes the technical strategy to manage the multi-disciplinary integration of hardware, embedded firmware, and cloud platforms as a sole engineer, ensuring system reliability and safety-critical execution.

1.1 Technical Project Control & Tools

- Configuration Management:

The GitHub repository serves as the Single Source of Truth (SSOT). To maintain strict self-discipline during independent development, interface data structures (JSON Payloads) and communication protocols are frozen as a personal baseline. This serves as a vital self-reminder: network ports (e.g., WebSocket port 8000 for Flutter Web, TCP port 1883 for the Python backend core) should only be modified after updating this blueprint first to ensure systematic cross-system consistency.

- Technical Performance Measures (TPMs):

  - Critical Payload Weight: The total weight of the smart payload box must not exceed 500 grams (housing the automatic EpiPen and the chemical oxygen candle).

  - Standard Response Time: Autonomous takeoff and aerial arrival at the incident coordinates must take less than 5 minutes.

  - Environmental Resilience: The VTOL drone architecture must maintain stable navigation under mountain wind speeds of up to 40 km/h.

1.2 Engineering Domain Integration & WBS (Work Breakdown Structure)

To manage system complexity independently, I have broken down the development lifecycle into three distinct technical domains that interface seamlessly through a predefined Interface Control Document (ICD):

- 1. Cloud & Application Communication Domain: Focuses on architecting the real-time data link between the application interface and the emergency cloud infrastructure. The ICD strictly defines this interface by mapping how the application publishes emergency alerts and subscribes to live telemetry through the cloud MQTT broker, freezing the exact JSON payload structures, topics, and network ports to ensure seamless, low-latency communication.

- 2. Hardware & Firmware Domain: Focuses on programming the ESP32 microcontroller in C++ to process real-time IMU sensor streams for automated fall-detection logic. It implements a non-connecting BLE Advertising protocol (Beacon Mode) to continuously broadcast vital telemetry directly inside the advertisement packets, successfully eliminating any user pairing requirements and bypassing standard control-center Bluetooth toggle restrictions, while utilizing peer-to-peer ESP-NOW protocols for zero-config proximity payload authentication.

- Autopilot & Edge AI Domain: Focuses on configuring the PX4 Software In The Loop (SITL) autonomous flight environment, developing the Python coordinate injection pipeline, and deploying the YOLOv8-Thermal computer vision model on the Raspberry Pi 5 edge computer.

1.3 Technical Risk & Mitigation Matrix

To ensure absolute system availability during critical emergencies, I have designed a fail-safe hierarchy that eliminates any single point of failure (SPOF):

| Risk ID | Technical Risk | Impact | Preventative Engineering Strategy & Backup |
| --- | --- | --- | --- |
| RSK-01 | drop in 5G/4G network in mountain. | Critical | switch to LoRa backup to broadcast coordinate |
| RSK-02 | Autonomous flight mission rejection by the PX4 Autopilot system during coordinate injection. | High | Programmatically enforcing strict landing waypoint parameters within the Python coordinate injection script to guarantee immediate mission validation and execution. |
| RSK-03 | Bystander or patient ignorance regarding the proper deployment and utilization of the delivered medical payload (EpiPen/Oxygen). | High | Integration of a rugged, embedded display on the smart payload box coupled with a live, bidirectional WebRTC audio/video broadcast connecting the emergency physician directly to the scene to visually guide the user. |
| RSK-04 | Delivery or unlocking of an incompatible or hazardous medical payload due to lack of patient clinical history verification. | Critical | Architecture of a cloud-native Medical Record Verification System that automatically cross-references the patient's authenticated electronic health records (EHR) before authorizing the drone dispatch and payload lock release. |

2. Concept of Operations (ConOps)

This section outlines the operational framework, stakeholder roles, and deployment lifecycle of the AeroMedic system. It defines how the system functions in real-world environments from initial onboarding to emergency execution.

2.1 Operational Stakeholders & User Roles

To ensure seamless execution, the system architecture maps to three core operational stakeholders:

- The Patient: Wears the smart wristband, which locally monitors vital signs and fall-detection telemetry.

- The Bystander / Responder: The individual physically present at the emergency scene who interacts with the arrived drone payload box, following visual displays and live medical audio to administer the critical payload.

- The Emergency Physician: Accesses the centralized cloud management dashboard to monitor critical telemetry incoming during an active alert and remotely oversees the autonomous drone dispatch logic.

2.2 Core Operational Lifecycle Scenarios

Scenario A: Device Onboarding and Lifecycle Maintenance (The Binding Flow)

- Patient Registration: The patient enters clinical profiles and emergency contact details through the application interface, creating an entry in the cloud database where the wearable device field remains empty.

- Hardware Authentication: The patient powers on the smart wristband, which begins broadcasting its unique hardware ID via non-connecting BLE advertisement packets.

- Proximity Binding: The mobile application detects the nearby advertisement packet, captures the hardware identifier, and patches the cloud database to bind the specific wristband to the patient record once.

- Predictive Maintenance: the wristband  wakes up intermittently (every 12 hours or when battery drops below 20%) to publish a brief maintenance packet to the cloud backend, enabling proactive device health tracking and failure prediction.

Scenario B: Critical Emergency Detection and Autonomous Dispatch

- Incident Trigger: The wristband detects a critical health threshold or a severe fall using embedded IMU streams, passing an alert packet locally to the application, or via long-range LoRa links if cellular infrastructure is missing.

- Cloud Orchestration & Verification: The cloud infrastructure intercepts the alert, automatically cross-references the patient's authenticated electronic health records (EHR) to verify payload compatibility, and approves dispatch.

- Flight Execution: The system injects the target coordinate parameters into the autopilot environment, launching the autonomous drone to deliver the payload box to the precise emergency site within 5 minutes.

- Scene Integration: The container unlocks via proximity ESP-NOW hardware authentication or a physician's remote cloud command. As an absolute fail-safe, if all network communications drop and the system becomes completely unresponsive, the container triggers an automatic timeout unlock to guarantee medical access. Simultaneously, a WebRTC audio/video link is launched to assist the bystander.

2.3 Operational Constraints

- Regulatory Airspace: The drone dispatch sequence must comply with local civilian aviation regulations and respect strict no-fly zones mapped in the autopilot mission parameters.

- Power & Logistics: The smart payload box must remain physically locked and low-powered until the autonomous drone confirms a safe arrival at the landing waypoint.

2.4 Functional Emergency Response Pathways (Modes of Activation)

To guarantee system availability under all environmental and user conditions, the architecture supports four distinct multi-layered emergency response pathways. Each path cuts through the functional layers from source to execution as detailed below:

Path 1: Automatic Wearable Device Pathway (The Primary Autonomous Link)

- Source Layer: The smart wristband automatically detects a critical biomedical anomaly or fall event.

- Transport Layer: Telemetry is transmitted locally via BLE to the patient's Mobile Application, which appends precise GPS coordinates and network timestamps, pushing the bundle via 5G/4G cellular links.

- Cloud & Logic Layer: The centralized Cloud core performs a fail-safe check, queries the nearest available drone station, and issues a structured Takeoff Order.

- Execution Layer: The targeted Drone Station processes the API Takeoff Order, launching the autonomous drone to deliver the medical payload to the location while providing continuous monitoring.

Path 2: LoRa Fail-Safe Wearable Device Pathway

- Source Layer: The smart wristband triggers an emergency alert in a rugged or remote mountain region lacking cellular coverage (or if the App didn’t send an ack).

- Transport Layer: The wristband broadcasts compressed emergency packets over long distances via LoRa directly to a local AeroMedic Ground Station Gateway.

- Cloud & Logic Layer: The Cloud core ingests the ground station payload, validates the incident location, and computes flight trajectories.

- Execution Layer: The drone is dispatched immediately, operating on a hardened communication loop to ensure secure delivery.

Path 3: Manual User App Pathway (The Direct On-Demand Link)

- Source Layer: The patient remains conscious but distressed and manually triggers the panic command via the application UI.

- Transport Layer: The request is routed instantly through 5G/4G protocols, attaching real-time mobile GPS positioning alongside the patient’s stored core profile data.

- Cloud & Logic Layer: The Cloud orchestrator bypassing automated sensor evaluation, immediately proceeding to flight clearance and station allocation.

- Execution Layer: The Takeoff Order is dispatched to the drone station for immediate launch and payload distribution.

Path 4: Voice Call User Pathway (The Legacy Fail-Safe Link)

- Source Layer: A bystander with no technical onboarding or access to the AeroMedic application encounters the incident and initiates a standard voice call via phone.

- Transport Layer: The emergency dispatcher intercepts the voice call over regular telecom networks and inputs the verbal location and patient details into the central Emergency Application platform, routing it via 5G/4G.

- Cloud & Logic Layer: The Cloud backend ingests the manually entered dispatcher parameters, registers it as an verified emergency, and schedules the flight logic.

- Execution Layer: The system triggers the drone station takeoff order, bringing the critical payload to the caller's verified location.

3 System-Level Requirements

- Functional Requirements (FR)

| ID | Requirement Description | Priority |
| --- | --- | --- |
| FR-01 | The smart wristband shall detect sudden fall events utilizing an on-board Inertial Measurement Unit (IMU) and generate an immediate emergency alert signal. | Critical |
| FR-02 | The system shall support four redundant communication channels for alert reporting. | Critical |
| FR-03 | The cloud platform shall automatically verify the patient's Electronic Health Record (EHR) to confirm payload suitability before approving the drone deployment. | High |
| FR-04 | The drone navigation system shall receive target coordinates, initiate autonomous takeoff, and navigate to the incident location without human intervention. | Critical |
| FR-05 | The smart payload box shall trigger an electronic locking mechanism to unlock automatically upon arrival confirmation via (ESP-NOW, cloud command, or fallback timer). | Critical |
| FR-06 | The system shall start and manage a audio/video streaming session using WebRTC between the on-duty physician and the first responder. | High |
| FR-07 | The smart wristband shall transmit periodic heartbeat signals to the cloud platform to monitor battery health and network connection status. | Medium |
| FR-08 | The system shall detect and bind the wristband to the proximity patient using BLE Beacon broadcasting without requiring a manual pairing process. | High |

- Non-Functional Requirements (NFR)

| ID | Quality Attribute | Strict Metric / Measurement Criteria |
| --- | --- | --- |
| NFR-01 | Alert Processing Latency | The duration from fall detection to alert transmission to the cloud shall be <3 seconds. |
| NFR-02 | Dispatch Response Time | The time from receiving medical cloud clearance to actual drone takeoff shall be <30 seconds. |
| NFR-03 | Mission Flight Time | The maximum duration for the VTOL drone to reach the furthest point within the operational boundary shall be <5 minutes. |
| NFR-04 | Payload Weight Limitation | The total weight of the smart payload box (loaded with the EpiPen) shall be ≤500 grams. |
| NFR-05 | Wind Resistance | The VTOL drone aerodynamic structure shall withstand crosswinds and gusts up to 40 km/h. |
| NFR-06 | Network Availability | High-availability fallback routing between 5G and LoRa networks shall guarantee a connection uptime of ≥99.9%. |
| NFR-07 | Wristband Battery Life | The smart wristband power management system shall sustain continuous operational broadcasting for ≥12 hours per charge. |
| NFR-08 | Geospatial Accuracy | The GPS/GNSS positioning subsystem for both the drone and the wristband shall maintain an accuracy radius within ±2.5 meters. |
| NFR-09 | Data Security & Privacy | All telemetry and patient data transmitted between edge nodes and the cloud shall be encrypted using TLS 1.3 and AES-256. |
| NFR-10 | Thermal Protection | The payload enclosure insulation shall maintain an internal temperature between 15∘C and 30∘C during flight to preserve medication viability. |

4. Subsystem Requirements Specification

This section defines the detailed technical requirements for each subsystem within the AeroMedic system. The requirements are structured following established systems engineering practices, with software-related requirements aligned to the IEEE 830-1998 standard and the ISO/IEC 25010 quality model where applicable.

4.1 Wearable Subsystem Requirements

4.1.1 Overview

The Wearable Subsystem is the primary patient-side component of the AeroMedic system. It is responsible for continuous monitoring of vital signs (SpO₂ and heart rate), fall detection, and emergency alert initiation. The device operates as a low-power, always-on health monitor that communicates with both the mobile application (via BLE) and the Drone Station (via LoRa as a fail-safe),and the Cloude using Wi-Fi.

Key Components:

- ESP32-PICO Microcontroller

- MAX30102 Pulse Oximeter & Heart Rate Sensor

- 6-Axis IMU (Accelerometer + Gyroscope)

- BLE 5.0 Module

- LoRa Transceiver (SX1276)

- Rechargeable Li-Po Battery (≥ 12 hours)

4.1.2 Functional Requirements (FR)

| ID | Requirement | Priority |
| --- | --- | --- |
| W-FR-01 | The wearable device shall continuously monitor the patient's SpO₂ (blood oxygen saturation) and heart rate at intervals ≤ 30 seconds. | Critical |
| W-FR-02 | The wearable device shall detect fall events using IMU data analysis within < 3 seconds of the event occurrence. | Critical |
| W-FR-03 | The wearable device shall transmit emergency alerts to the mobile application via BLE within < 1 second of detecting a critical event. | Critical |
| W-FR-04 | The wearable device shall broadcast its unique hardware identifier via BLE advertisement packets (Beacon Mode) for proximity-based patient binding, eliminating the need for manual pairing. | High |
| W-FR-05 | The wearable device shall utilize BLE as the primary communication channel, activating Wi-Fi as a secondary fail-safe only when BLE connectivity to the application is lost for > 10 s, and LoRa for emergency's. | Critical |
| W-FR-06 | The wearable device shall transmit periodic health status packets to the cloud backend every 12 hours for predictive maintenance and battery monitoring. | Medium |
| W-FR-07 | The wearable device shall store emergency alert data locally (in flash memory) for up to 24 hours if no network connectivity is available, and retransmit when connectivity is restored. | High |
| W-FR-08 | The wearable device shall provide visual (LED) and haptic (vibration) feedback to the patient upon successful alert transmission or when anomaly happened and when the system acknowledges receipt. | Medium |
| W-FR-09 | The wearable device shall keep the Wi-Fi module in an ultra-low-power Deep Sleep state by default, power it on only during BLE failures to transmit messages and shut it down immediately after transmission acknowledgment. | High |
| W-FR-9.5 | The wearable device shall compile a daily health summary packet (including average/peak heart rate, SpO₂ trends, and sleep/activity duration) and transmit it to the cloud via Wi-Fi or BLE once every 24 hours to assist in long-term medical diagnosis. | High |

4.1.3 Non-Functional Requirements (NFR)

- Performance Requirements

| ID | Requirement | Target |
| --- | --- | --- |
| W-NFR-01 | Fall detection latency from event occurrence to alert generation | < 3 seconds |
| W-NFR-02 | BLE alert transmission latency from detection to mobile app receipt | < 1 second |
| W-NFR-03 | Device boot-up time from power-on to full operation | < 5 seconds |
| W-NFR-04 | Sensor data acquisition and processing cycle time | ≤ 5 seconds |

- Power & Battery Requirements

| ID | Requirement | Target |
| --- | --- | --- |
| W-NFR-05 | Continuous battery life under normal monitoring conditions | ≥ 12 hours |
| W-NFR-06 | Battery charging time (0% to 80%) via USB-C | ≤ 2 hours |

- Environmental Requirements

| ID | Requirement |
| --- | --- |
| W-NFR-07 | The wearable device shall be waterproof with an IP67 rating (dust-tight and protected against immersion up to 1 meter). |
| W-NFR-8 | The wearable device shall operate within temperature range: 0°C to 80°C. |
| W-NFR-9 | The wearable device shall withstand minor impacts and drops (up to 1.5 meters). |
| W-NFR-10 | The device shall be lightweight, with total weight ≤ 100 grams. |

- Communication & Interface Requirements

| ID | Requirement | Interface |
| --- | --- | --- |
| W-NFR-11 | The wearable device shall continuously stream vital signs and alerts via BLE to the mobile application under normal operating conditions. | BLE |
| W-NFR-12 | The wearable device shall utilize 2.4 GHz Wi-Fi to establish a direct-to-cloud connection only when the mobile application is unresponsive and a pre-configured Wi-Fi network is available. | Wi-Fi |
| W-NFR-13 | The wearable device shall communicate with the Drone station via LoRa (SX1276) at with a range ≥ 15 km in open terrain, as a fail-safe when cellular networks are unavailable. | LoRa |
| W-NFR-14 | All data transmitted over BLE shall be encrypted using AES-128 encryption. | BLE |
| W-NFR-15 | All data transmitted over LoRa shall be encrypted using AES-128 encryption. | LoRa |
| W-NFR-16 | All data transmitted over Wi-Fi shall be secured using TLS 1.2 (HTTPS/MQTT over TLS) to ensure secure end-to-end cloud ingestion. | Wi-Fi |
| W-NFR-16.5 | The wearable device shall authenticate with the cloud backend using device-specific X.509 certificates. | Wi-Fi |

- Security, Safety & Privacy Requirements

| ID | Requirement |
| --- | --- |
| W-NFR-17 | The wearable device shall store patient health data locally in encrypted format using AES-128. |
| W-NFR-18 | The device shall implement a factory reset mechanism that permanently erases all stored patient data. |
| W-NFR-19 | The device shall include tamper detection mechanisms that lock the device and erase data upon unauthorized physical access attempts. |
| W-NFR-20 | The device shall only transmit anonymized patient identifiers (not full personal information) over wireless networks. |

- Reliability & Maintainability Requirements

| ID | Requirement |
| --- | --- |
| W-NFR-21 | The wearable device shall have a Mean Time Between Failures (MTBF) of ≥ 10,000 hours. |
| W-NFR-22 | The device shall support firmware updates over-the-air (OTA) via BLE or LoRa (upon physical proximity). |
| W-NFR-23 | The device shall record diagnostic logs for the last 30 days of operation, including battery health, sensor errors, and communication attempts. |

- Usability Requirements

| ID | Requirement |
| --- | --- |
| W-NFR-24 | The wearable device shall be designed for single-handed operation and comfortable continuous wear. |
| W-NFR-25 | The device shall include a physical panic button for manual emergency activation that is easily accessible to the patient. |

4.2 Mobile Application Subsystem Requirements

4.2.1 Overview

The Mobile Application Subsystem acts as a critical communication bridge within the AeroMedic ecosystem. Developed using Flutter and Dart, this cross-platform application links the patient's bio-wearable device (via BLE) to the cloud backend (via MQTTS) and tracks the autonomous medical drone dispatch. The application is designed to prioritize ultra-low latency, strict data privacy, high visual contrast under stress, and robust operating system persistence.

Key Architectural Components:

- Framework & Language: Flutter & Dart (Clean Architecture with BLoC/Provider state management).

- Local Communication Interface: flutter_blue_plus for BLE data ingestion.

- Cloud Telemetry Client: MQTTS Client with mutual TLS (mTLS) authentication.

- 4.2.2 Functional Requirements (FR)

| ID | Requirement | Priority |
| --- | --- | --- |
| APP-FR-01 | The application shall allow users to create and securely store a local medical profile including critical allergies, blood type, and respiratory history. | High |
| APP-FR-02 | The application shall continuously scan for and identify the unique hardware identifier of the AeroMedic Wearable Subsystem via BLE advertisement packets (Beacon Mode). | Critical |
| APP-FR-03 | The application shall intercept and process the live stream of SpO2 and heart rate telemetry embedded directly within the wearable's broadcasted BLE beacon payloads. | Critical |
| APP-FR-04 | The application shall execute a 20-second visual and auditory pre-alert countdown upon receiving an emergency trigger (SpO2<85% accompanied by positive IMU fall detection data). | Critical |
| APP-FR-05 | The application shall present a prominent, high-visibility "Cancel Alert" button alongside a secondary "Force Send Now" bypass button during the 20-second countdown. | Critical |
| APP-FR-06 | The application shall automatically compile, encrypt, and publish an emergency payload (including GPS coordinates, timestamp, and medical profile) to the cloud via MQTT if the 20-second countdown expires or when "Force Send" Butten pressed. | Critical |
| APP-FR-07 | The application shall display early warning push notifications and haptic vibrations when a gradual decline in oxygen saturation (SpO2<90%) is registered. | Medium |
| APP-FR-08 | The application shall provide a manual Trigger on the primary dashboard for immediate drone requests. | High |
| APP-FR-09 | The application shall render a map showing the drone's live route, current coordinates, and Estimated Time of Arrival (ETA). | High |
| APP-FR-10 | The application shall compute the distance to the landed drone via BLE RSSI/UWB, and automatically transmit a secure proximity token to release the payload box's lock when the user is within < 2 meters. | Critical |
| APP-FR-11 | The application shall utilize native Android (SYSTEM_ALERT_WINDOW) and iOS (Critical Alerts) APIs to bypass system locks, mute switches, or "Do Not Disturb" modes during an emergency countdown. | Critical |
| APP-FR-12 | The application shall programmatically initiate a direct cellular voice call to the Saudi Red Cross (997) in parallel with the MQTT cloud alert transmission upon countdown expiration. | Critical |
| APP-FR-12.2 | The application shall monitor the smartphone's native battery level (%) and low-power mode status via system APIs in real-time. | High |
| APP-FR-12.8 | In Eco-Emergency Mode, the application shall suspend non-critical background services (including map tile downloads, UI animations, and log uploads), retaining only BLE ingestion, GPS acquisition, and MQTT publishing. | Critical |

- 4.2.3 Non-Functional Requirements (NFR)

- Performance & Response Time Requirements

| ID | Requirement | Target |
| --- | --- | --- |
| APP-NFR-01 | Processing and transmission latency of the emergency payload to the MQTT broker upon countdown expiration | < 1.5 seconds |
| APP-NFR-02 | BLE data packet ingestion and UI state refresh rate | ≤ 500 ms |
| APP-NFR-03 | Google Maps asset position rendering refresh interval | ≤ 1 second |

- Security, Safety & Privacy Requirements

| ID | Requirement |
| --- | --- |
| APP-NFR-04 | All outgoing telemetry, vital signs, and patient health profiles must be encrypted in transit using MQTTS (MQTT over TLS 1.3) with mutual authentication (mTLS). |
| APP-NFR-05 | Local caching of sensitive medical profiles and authentication tokens must be encrypted using AES-256 inside secure OS sandboxes (Keychain/Keystore). |
| APP-NFR-06 | The application must enforce strict data minimization, ensuring personal identifiers are masked or tokenized before cloud ingestion, adhering to local healthcare privacy frameworks. |

- Reliability, Architecture & Maintainability Requirements

| ID | Requirement |
| --- | --- |
| APP-NFR-07 | The codebase must strictly adhere to Clean Architecture principles, ensuring that data layers, domain entities, and presentation presentation logic are decoupled to support future sensor additions. |
| APP-NFR-08 | The application must achieve a crash-free session rate of ≥ 99.9% across both supported platforms (iOS and Android). |

- Usability Under Stress Requirements

| ID | Requirement |
| --- | --- |
| APP-NFR-09 | The UI must support a high-contrast theme with large, touch-target buttons (minimum 64x64 density pixels) usable under direct sunlight and high-stress scenarios. |
| APP-NFR-10 | The layout must prevent any screen dimming or automatic sleep timeouts while an active emergency event or countdown is underway. |

- OS Permissions & Background Management Requirements

| ID | Requirement |
| --- | --- |
| APP-NFR-11 | The BLE background service must run as an OS-level Foreground Service (with persistent notification) to prevent the operating system from terminating the process to save memory. |
| APP-NFR-12 | The application must gracefully re-acquire or prompt for critical permissions (Location, Bluetooth, Background Fetch, Overlay) upon startup if revoked by the user. |

- Network Resilience & Interoperability Requirements

| ID | Requirement |
| --- | --- |
| APP-NFR-13 | In the event of 4G/5G connection loss during an emergency trigger, the app must queue the data packet locally and re-attempt transmission at 1-second intervals until a handshake is confirmed. |
| APP-NFR-14 | The application shall implement an automatic data-saving mode under highly degraded networks (Edge/3G), suspending non-essential map tile downloads and prioritizing the ingestion of raw, low-bandwidth coordinate strings (lat/long) to sustain live drone tracking. |

- State Management & Concurrency Requirements

| ID | Requirement |
| --- | --- |
| APP-NFR-15 | The application must run asynchronous background isolates for BLE incoming streams to prevent high-frequency sensor data from locking or lagging the UI thread. |
| APP-NFR-16 | The state management layer (BLoC/Provider) must guarantee that the "Cancel Alert" button interaction takes absolute event-loop priority over telemetry updates. |